Initiating New project:

A comprehensive look at the two widely used sedimentation velocity software, SEDFIT and DCDT+ :

To study the proper way of doing fit using pure CRP in TEK100 buffer as raw sedimentation profiles for data analysis.

Background:

Modern computational technique has led to considerable improvements in the determination of weight-average sedimentation coefficient via differential sedimentation coefficient distributions, which have the potential to discriminate different sedimentation species.  In 1992, Stafford has shown how an apparent sedimentation coefficient distribution g(s*) can be calculate from a transformation of the time-derivative of the sedimentation profiles. John philo was able to compile this into window operating system and make it a user friendly Software calls DCDT+.

More recently, Peter Schuck from NIH  showed how an apparent sedimentation coefficient distribution Is-g*(S) can be calculate directly from least-squire(Is) modeling of the sedimentation profile, permitting higher precision through the use of an increased data basis and permitting wider distributions and more general application. A sedimentation coefficient distributions c(s) with significantly higher resolution can be achieved through direct modeling and deconvolution of diffusional broadening of a complete set of sedimentation profiles. Thus SEDFIT software was created. Later SEPHAT software was introduced for Global modeling.  Best of all it is all free from the internet.

Goals of project:

To me fitting is an art form of science. In order to get a good fit with high precision, one needs to understand the important of scan choosing in relationship to the fitting profile and ask the question:  What is the criteria for choosing the correct number of scans for data analysis in order to obtain a good fit.

Preliminary Experiment Results and conclusion:

Using native CRP protein in Tek100 buffer as example.

In SEDFIT study, I divide the experiment into two sets.  Set 1, I start the analysis from scan 1.  Set 2,  I choose to start the analysis from scan 5 when the lower plateau is formed.

In DCDT+ study, I used scan that yields a good lower and upper plateau.

SEDFIT study result:

Set 1:

Set2:

Set 1:Using scan without lower plateau (start analysis from scan 1 up)

| Trial | Scan used | S20,w | Fit for one species |
| --- | --- | --- | --- |
| 1 | 1-100 | 3.523 | Good- |
| 2 | 1-40 | 3.530 | Good+ |
| 3 | 1-30 | 3.530 | Good + |
| 4 | 1-20 | 3.530 | Good+ |
| 5 | 1-15 | 3.521 | Good- |
| 6 | 1-10 | 3.523 | Poor |
| 7 | 1-8 | 3.541 | Poor |

Set 2:Exclude scan without lower plateau (start analysis from scan 5 up)

| Trial | Scan used | S20,w | Fit for one species |
| --- | --- | --- | --- |
| 1 | 5-100 | 3.511 | Good- |
| 2 | 5-40 | 3.520 | Good |
| 3 | 5-30 | 3.525 | Good+ |
| 4 | 5-20 | 3.520 | Good+ |
| 5 | 5-15 | 3.505 | Good |
| 6 | 5-10 | 3.418 | Poor |
| 7 | 5-8 | 3.252 | Poor |

The conclusion from this SEDFIT study shows that :

- Mininum scans required is 15scans or more

- There is no need to wait for lower plateau to form (can start scan from 1 up).

- It get better fit (sharper peak) if the upper plateau is begin to disappear(1-30, green, 1-40 magenta)

- More scans is ok but it does not improve the fit (scan 1-100, red)

- Not enough scan is bad because not only the S20, w peak is shifted but also the peak area(C) is affected (low).

- So to get a good fit in here, use 20-40 scans and include scan from scan 1.

DCDT+ study result:

| Trial | Scan used | S20,w | Fit for one species |
| --- | --- | --- | --- |
| 1 | 5-105 | - | - |
| 2 | 5-45 | 3.545 | Poor --- |
| 3 | 5-35 | 3.429 | Poor-- |
| 4 | 5-25 | 3.401 | Poor- |
| 5 | 5-15 | 3.451 | Fair+ |
| 6 | 5-10 | 3.515 | Fair++ |
| 7 | 5-8 (the minimum scan required) | 3.578 | Fair+++ |

The conclusion from this DCDT+ study shows that :

- Too little scan (minimum scan required by the software is 4 scans) will have a poor resolution but better fit (Trial 7).

- Too many scan without an upper plateau will not get a peak to fit (Trial 1).

- As a general rule. have a good lower plateau and upper plateau and reasonable amount of scan is important to achieve good fit with good resolution in DCDT+ analysis.

On the final analysis:

Given the best shot and compare these two softwares SEDFIT and DCDT+ (choosing the scan correctly, etc) we have:

Fit for one species

Fit for two species with M-D constraint:

Fitting for two species without constraint:

Final Conclusion:

- For SEDFIT, S20,w =3.530, fitting is good. Did not detect any aggregate or dissociate species.

- For DCDT+, S20,w =3.521, fitting is poor, but the s value is almost the same as the SEDFIT s value. But if we fit for two species, S20,w1 = 2.830, S20,w2 = 4.332. The fitting is good.  Or if we fit for two species with M-D constraint, S20,w1 = 2.831, S20,w2 =4.25. The fitting is also good. However the S20,w = 2.8 or 4.2 are way off for the known S20,w value for CRP.

- When we look at these results; the SEDFIT is right on the mark and the fitting is good. Not so for DCDT+, you have to choose between good number with poor fit or good fit with s value that is not quite accurate.

8888888888888888888888888888888888888888888888888888888888888888888888888888888888888

P.S. If interested, all the original raw fitting data are posted below for close scrutiny.

All SEDFIT Data analysis

Raw data taken from 114931, using SedFit software to do analysis;

Set 1: Using scans 1-100 (without waiting for the lower plateau to form)

Using scans 1-40

Using scans 1-30

Using scans 1-20

Using scans 1-15

Using scans 1-10

Using scans 1-8

Using scans 1-5

Using scans 1-4

$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$$

Set 2: Waiting for the lower plateau to form ;

Using scans 5-100 (scan 5 has a good lower plateau)

Using scans 5-40 (waiting for the lower plateau to form)

Using scans 5-30 (waiting for the lower plateau to form)

Using scans 5-20 (waiting for the lower plateau to form)

Using scans 5-15 (waiting for the lower plateau to form)

Using scans 5-10 (waiting for the lower plateau to form)

Using scans 5-8 (waiting for the lower plateau to form)

8888888888888888888888888888888888888888888888888888888888888

All  DCDT+  Data analysis

Same raw data taken from 114931, using DCDT+ software to do analysis:

Using scan 5-105

Using scan 5-45

Using scan 5-35

Using scan 5-25

Using scan 5-15

Using scan 5-10

Using scan 5-8