Light Intensity Variation Check

The typical light intensity variation of both the sample and reference sector across the cell is about 20%.  The important point is that both the sample and reference sector intensity at a given radial distance is equal.  If they are not nearly equal at a given radius, the absorbance output (O.D.) will be incorrect.

Checking light intensity variation across a cell

| Step | Action |
| --- | --- |
| 1. | Ensure power is on to the instrument. GUI is running. Rotor chamber is empty. |
| 2. | Install a windowless double sector cell assembly into rotor hole #2 or #4 depending on the rotor used. Install the counterbalance in #4 or #8. |
| 3. | Open the door. Install the rotor. Install the monochromator. Close the door |
| 4. | On the GUI menu bar go to File>New. The setup window for a new scan file will open. |

## Note

You can rename this file and save it for future reference to monitor instrument performance.

Checking light intensity variation across a cell (continued)

| Step | Action |
| --- | --- |
| 5. | Enter the rotor type. Set XL settings for the following: Speed = 3000, Time = hold, Temp = 25. |
| 6. | Enter the following parameters for cell #1: Select Absorbance, Select Velocity, Rmin = 5.8, Rmax = 7.3, W1 = 400, Comment = Intensity Variation. |
| 7. | Click on the Detail button and set the following: Radial Step Size = .001, Replicates = 3, Mode = Continuous. Click “OK”. The velocity window will close and you will be returned to the main scan setup window. |

Checking light intensity variation across a cell (continued)

| Step | Action |
| --- | --- |
| 7. | Click the Options button on the main scan setup window. Select the following: “Acquire Intensity data instead of OD data”, “Radial Calibration” before first scan. |
| 8. | Click the “Start Single Scan” button. When the scan is complete the intensity graph will be display on the screen. |
| 9. | Minimize the GUI. Import the scan file into Origin. a. On the Origin menu bar select File>Import>Single ASCII. A directory path dialog box will open. Select the path to the completed scan. Click “OK”. |
| 10. | On the Origin menu bar select Plot>Line. The plot of the graph will be displayed. |
| 11. | On the Origin tool bar select the cursor and move to the maximum intensity peak. Record the value. Next move the cursor to the minimum intensity peak. There should be a distance of 1 cm between peaks selected. Record the value. Use this formula to calculate the variation across the cell: % Variation = Max Intensity – Min Intensity X 100 Max Intensity This value should not exceed 20%. If the variation exceeds 20% then perform the strobe lamp adjustment and/or radial calibration. Verify the optics are clean. |
| 12. | If the variation is satisfactory then the check is complete. Stop the run. Close the scan file. Open the door. Remove the monochromator and rotor. Close the door. |

The strobe lamp should be cleaned if intensity is low, each PMC or as necessary.

Cleaning the strobe lamp

| Step | Action |
| --- | --- |
| 1. | Ensure power is off to the instrument. |
| 2. | Disconnect the lamp assembly from the lamp power supply. |
| 3. | Loosen the setscrew holding the lamp in the wavelength selector housing. |
| 4. | Remove the lamp housing assembly. Loosen the lamp adapter setscrew. Unscrew the lamp adapter. |
| 5. | Remove the o-ring on the lamp face. Set to the side or replace as necessary when cleaning is complete. Keep the lamp installed in the lamp housing. |
| 6. | Clean the strobe lamp face using standard white toothpaste and/or a soft pencil eraser. Wipe clean with distilled water soaked soft tissue paper. Dry with air. |
| 7. | Reinstall the lamp adapter. Tightened the lamp adapter setscrew. |
| 8. | Reinstall the lamp housing assembly into the wavelength selector housing. Tighten the setscrew. |
| 9. | Perform the necessary quality checks to verify lamp performance. |